TITAN BED

SAFETY

IMPORTANT SAFETY INFORMATION

1.  Please read the entire manual before operating the bed.  Failure to read

the manual can cause injury or undue stress on the equipment or yourself.

2.  Do not exceed the 750lb. weight capacity of the bed frame.

3.  Insure that the side rails are in the upright position when the patient is

in the bed.

4.  Run the power supply cord under the bed to insure safety from

running over it or accidentally unplugging the unit.

5.  In case of emergency the bed is equip with a red C.P.R. release handle

at the head of the bed.

6.  When the bed is in use make sure the casters are locked in place.

7.  Make sure the head of the bed frame is not touching the wall to insure

there is no damage to the wall or the bed frame.

8.  When transporting a patient into the bed make sure that the bed frame

is at its lowest position for an easy transport.

9.  When restraints are used, be sure to attach to some part of the bed that

articulates when the gatches are raised and lowered.  Failure to do so

could result in the restraints injuring the patient.

1

OPERATIONAL

INSTRUCTIONS

Bed (hi/low) function- press and hold the up or down buttons on the

handset with the picture of the bed.  This will raise and lower bed.

Head (up/down) function- press the top buttons on the handset with the picture of the head up or down.

Foot (up/down) function- press the bottom buttons on the handset with the picture of the foot up or down.

Trendelenburg function- press the buttons on the handset with the picture of the bed in the trendelenburg position.

Operating instructions continued >	2

PATIENT RESTRAINTS

If patient restraints are used, they must be attached to some part of the bed that articulates when the gatches are raised and lowered.  Failure to do so could result in the restraints injuring the patient.

PATIENT TRANSFERS

1)  Set the brake and steer on all wheels.

2)  Adjust the bed to match the transfer height (it is always better to transfer down).

3)  Lower the side rails or remove them.

4)  Use a hospital sheet or draw sheet and appropriate number of personnel to lift the patient on/off the bed.

TRANSPORTING THE BED FRAME-

1)  Make sure you have ONE of the two steer wheels locked in place.  Now you can push the bed from the opposite end and use the steer wheel to insure a straight roll down the hall.

2)  Unplug the bed before transporting out of the room or down the hall

3)  For best stability lower the bed to the lowest position before rolling.

4)  Use a minimum of two persons when moving the bed for added stability

5)  USE CAUTION when moving the bed in/out of elevators, or when rolling on uneven surfaces.  Failure to use caution can result in injury to equipment or persons.

SIDE RAIL ADJUSTMENTS

To attach the side rails to the bed you must insert the slide rail frame into the horizontal tube under the bed.  Adjust in or out until desired placement and tighten with the black knob under the bed frame.  Adjust the vertical height with the black knob on the side of the hand rail assembly.

OPTIONAL PADDED SIDE RAILS

To drop the padded side rails pull up on the horizontal bar behind the side rail and swing up/down until the spring loaded pin stops in the desired placement.  To remove padded side rails loosen the black knob under the bed until slide bar moves freely.

3

WARRANTY

(5 YEAR)  METAL BED FRAME

(2 YEAR) LINAK MOTORS & MECHANICAL PARTS

(2 YEAR) FOAM MATTRESS

QUESTIONS OR COMMENTS

CALL US AT (281) 324-2900

FAX (281) 324-3090

4

TITAN

USER/PARTS

MANUAL

MODEL

3600E/750

TELE-MADE, INC.

3215 HUFFMAN EASTGATE

HUFFMAN, TX 77336

(281) 324-2900

FAX (281) 324-3090

ASSEMBLY INSTRUCTIONS

The TITAN Bariatric bed ships in 2 sections for ease in shipping and handling.

Step 1		Align main head section of frame with main foot section frame.

Step 2		2 carriage bolts 2 ¾” in length are used to secure the two frames

together on left and right side, allow weight of frame to relax

to floor in the middle of bed.  Tighten bolts using 9/16 wrench.

Step 3		Lift in mid section of frame, in order to start ¾” nuts to tension

bar.  Tighten ¾” nuts to tension bar evenly, left and right side.

It may be helpful to prop mid section of bed up while tightening

tension bar.  This will relieve pressure on bar.

Step 4		Plug in foot motor and foot hi/low motor, each should be color			coded.

Step 5		Plug in hand control.

Step 6		Add head and foot boards.

Step 7		Run bed through complete operational test.

5

TITAN BED

TECHNICAL SPECIFICATIONS

Total length							93”

Sleep surface							84”

Weight capacity						750 lb.

Bed deck low							15”

Bed deck high							28”

Caster size							5”

Total weight							475 lb.

Power requirements						110V AC 60Hertz

Motors								DC Volts

Battery backup						Yes

CLASSIFICATION AND STANDARDS

The TITAN Bariatric bed is designed and manufactured to meet the following equipment classifications and standards:

Technical and Quality Assurance				UL 2601-1

ETL Listed CSA

Equipment Classification per EN60601-1			Class 1 equipment,

internally powered										equipment.

Degree of protection against electric shock			Type B

Degree of protection against the presence of 			Not for use with flammable

flammable anesthetic mixtures				anesthetics.  Not to be used								with oxygen tents.  Use

oxygen administered

equipment of the nasal, mask,								or ventilator only.

6

TITAN BED

Model 3600E

PARTS LIST

DESCRIPTION							PART#

Side Rail								100SR

Slider for side rail							100SSR

Side rail adj. knob							100RAK

Base frame head end							100BFHE

Base frame foot end							100BFFE

Head deck								100HD ____width

Upper leg deck							100ULD ____width

Lower leg deck							100LLD  ____width

Transfer caster bar							100TCB

Tension bar								100TB

MOTOR ASSEMBLY

Head Motor								LA31 head

Foot Motor								LA31 foot

Hi/Low Motor								LA31 hi/low

Control Unit								CB9

Battery Backup							BA1810

Hand Control								100HC

7

TABLE OF CONTENTS

Important Safety Information						1

Operational Instructions							2

Patient Restraints								3

Patient Transfers								3

Transporting the Bed Frame						3

Side Rail Adjustments							3

Optional Padded Side Rails						3

Warranty									4

Assembly Instructions							5

Technical Specifications							6

Classification and Standards						6

Parts List									7

Motor Assembly								7

Major Components							8

# Text boxes

HEAD

HI/LOW

TREN.

KNEE